## 1
|  |  |  | 5000000 | 3257700 | 2696000 | 3,050,00.00 | 1378250 | 1850000 | 4694030 | 3350000 | 2275782 | 2180600 | 5907921
Item No | Specification | Requirement | Priority | Delmege | research instrument | P&T | Microtech | bhoomitech | Biomedite | Hemsons | premiuM | cmd | Analytical | avon
1 | Equipment | Refrigerated Microcentrifuge
| Make | (Please specify) |  | Thermo Scientific | Dlab | YINGTAI | ortoalresa | SciTEK | shuke | Ohaus | Hettich | Infitek | Sorvall legend Micro 21R | Eppendorf
| Model and Model year | (Please specify) |  | Multifuge X1R Pro/ 2021 (we provide machine manufactured in 2025) | DG1616R | TGL18 | Digicen 22R/2025 | CFG-TI5HR/2025 | TGL-1650,2025 | FC5718RIVD | Universal 320R | CFGR-B21B |  | 5804R/5805000327
| Country of origin | (Please specify) |  | Germany | China | CHINA | Spain | China | China | USA | Germany | China | China | Germany
| General | Table-top, refrigerated microcentrifuge for the intended use in routine diagnostic applications (PCR sample prep, RNA/DNA extraction, protein assays, etc.) |  | meet research and diagnostic purpose | yes | yes | yes |  |  | Yes | Yes
| Maximum Speed with the given rotor/ rpm | ≥ 15,000 rpm | C | 15200rpm with micoliter 30*2 | yes | yes | yes | 15000rpm | 16500 | Better, Upto 18,000rpm | 16,000rpm | 21000 RPM | 14800rpm | 14000rpm
| Maximum Speed with the given rotor/ RCF | ≥ 20,000 × g | C | 25830g with micoliter 30*2 | yes | yes | yes | 21380g | 24814g | Better, Upto 23,542 x g | 24,900g | 29828 g | 21100*g | yes, 20913g
| Speed Setting | Adjustable in increments of 100 rpm |  | 1rpm increment | yes | yes | yes | yes | not mentioned in the boucher | Better, 10 acceleration gear | Adjustable in 10 |  | - | yes
| Temperature Range | -10°C to +20°C or better |  | -10 to +40 | yes | yes | yes | yes | -20 to 40 | Better, -20C to 40C | -10C to 40C |  | yes | yes
| Cooling System | CFC-free refrigeration, fast pre-cooling (<10 min to 4°C) | C | cfc free R134a | yes | yes | yes | yes | not mentioned in the boucher | From room temperature to 4C <10min | Yes |  | yes | yes
| Temperature Accuracy | ±1°C | C | yes | yes | yes | yes | yes | yes | Yes | Yes |  | - | yes
| Rotor Type | Fixed-angle rotor (compatible with 1.5/2.0 mL microtubes) | C | yes | yes | yes | yes | yes, 2.0ml*24 |  | Yes | Provided cut # 1552 |  | yes | yes
| Rotor Capacity | Fixed-angle rotor ≥ 24 × 1.5/2.0 mL microtubes, adaptors for 0.2 mL and 0.5 mL microtubes | C | capasity- 30*1.5/2.0ml, Adapters given for 0.2ml and 0.5ml | No adapters rotor only | yes | yes | yes |  | Yes | 24 x 1.5/2.0ml, adapters for 0.2ml & 0.5ml is provided | 24*1.5/2ml rotor with 0.2/0.5ml adaptors | yes | yes
| Accessories | Swing rotor with the capacity of 10- 16 holes to allocate 5ml tubes ((LxØ): 75 x 13 mm) Maximum Speed/ rpm 4,000 rpm or better | C | yes, swing out rotor given with Adapter- 32*5/7ml, max speed-5500rpm | yes | yes | yes | 18*5ml angle rotor | 16*5ml aerosol tight angle rotor | Better rpm upto 5000rpm 16 tubes | 28 x 5ml tubes; max speed 5000rpm/4193xg | swingout 4*8*15ml with 5ml adaptors for blood collection tubes | - | yes, 4 bucket for 2*50ml conical tube. Accommodate tube from 0.2ml to 100ml
| Acceleration/Deceleration | At least 9 acceleration and 9 deceleration profiles Pulse free acceleration and deceleration | C | yes | yes | yes | yes | yes | No, 7 profiles | Yes | 9 acceleration and 9 deceleration including break free decelleration |  | - | yes
| Body | The centrifuge should be manufactured with quality materials providing broad chemical resistance with exceptional strength for durability and reliability | C | yes | yes | yes | yes | yes |  | Yes | SS chamber and metal housing and lid is available |  | - | yes
| Timer | 1 sec – 99 min with continuous run option | C | 9h,59min with continuous run option | yes | yes | yes | yes |  | Yes | Available |  | yes | yes
| Display | Digital LCD or LED display showing speed (rpm/RCF), time, and temperature | C | yes | yes | yes | yes | LED |  | Yes | Available | LCD | yes | yes
| Noise Level | ≤ 60 dB | C | yes | yes | yes | yes | yes |  | <= 60 +-2 dB (A) | 54dB with angle rotor |  | yes | yes
| Imbalance Detection | Automatic rotor imbalance detection | C | yes | yes | yes | yes | yes |  | Yes | Available |  | yes | yes
| Lid Lock | Automatic lid lock with safety interlock function |  | motorized lid lock | Manual lid closer | yes | yes | yes | no | Yes | Available |  | yes | yes
| Memory Function | Minimum 6 user-defined program storage or better |  | upto 100 | yes | yes | yes | 9 programs | 7 users | Yes, Better upto 99 user programmes | 9 programme memories | 6 No.S | yes | yes
| Power Supply | 220–240 V, 50 Hz (compatible with Sri Lankan standards) | C | yes | yes | yes | yes |  |  | Yes | Compatible |  | yes | yes
| Certifications for the equipment to be used in diagnostic settings | CE-IVD Certified Declaration and or NMRA certificate | C | Yes,CE-IVD, NMRA certificate | No NMRA registration or IVD | Q: Please re-attach CE-IVD , NMRA certificate (Mentioned as attached in the specification sheet but the attachment was not available) A: CE-IVD Certification- Attached [No. Provided certifications are not NMRA registration or IVD] | yes, CE-IVDR [Pre bid minute: we give priority to instrument with both NMRA and CE-IVD] | CE | NO | Q: Please re-attach CE-IVD , NMRA certificate (Mentioned as attached in the specification sheet but the attachment was not available) A: Please user the below link to access to the CE-IVD Certificated related to the quoted Refrigerated Microcentrifuge. https://ec.europa.eu/tools/eudamed/#/screen/search-device/37c7a72b-1ae4-426f-8310-97bbb448c512 [ No. CE certification does not include IVD] | Q: Please re-attach CE-IVD , NMRA certificate (Mentioned as attached in the specification sheet but the attachment was not available) A: Kindly find the attached EU quality management certificate for your reference. [[No. Provided certifications are not NMRA registration or IVD] | CE onLy | yes | CE,IVD yes, No NMRA
| Certifications for the Manufacturer | ISO 13485 | C | yes | yes | yes | yes | yes | yes | Yes | Available & submitted | yes | yes | yes
| UPS Power Backup | Online UPS Power back up for a minimum of 15 minutes | C | yes | yes | yes | yes |  | yes | Yes, Locally quoted | Supplied | yes | yes | yes
| Warranty | Local Agent should provide a warranty for the instrument including UPS system from the manufacturer for 2 years or more. | C | yes | yes | yes | yes | 2 years for the main unit, 1 year for UPS | yes | Yes | 2 years manufacturer warrant for manufacturing defects is provided. | yes | yes | yes, 3 years for the instrument 1 year UPS
| Maintenance | Maintenance of the equipment should be every 06 months free of charge during the warranty period | C | yes | yes | yes | yes | yes | yes | Yes | Provided | yes | yes | yes
| Availability of Spare parts | The Local agent should assure to make spare parts available up to 10 years |  | yes | yes | yes | yes | yes | yes | Yes | Provided | yes | yes | yes

## 2
[5 empty] | 593000 | 551500 | 845000 | 589000 | 785000 | 2332895 |  | 348550 | 518600 | 1383660
Item No | Specification | Requirement | Priority | Delmege | research instrument | P&T | Microtech | bhoomitech | Biomedite | Hemsons | premiuM | cmd | Analytical | avon
2 | Equipment | Forced Air-Drying Oven
| Make | (Specify) |  | NQ | Infitek | Bio Base | Boxun | SCITEK | China-Zenith LAB | Memmert | NQ | Infitek | biobase | LK lab
| Model | (Specify) |  | NQ | DOF-v140A | BOV-V225FI | BGZ-306 | FDO-240A | PPH-240A | UF260 | NQ | DOF-H250B | BJPX-HGZ270 | LO-FS250
| Country Of Manufacture | (Specify) |  | NQ | China | China | China | China | China | Germany | NQ | China | China | Korea
| Capacity (L) | 250L or better | C | NQ | 140L | 240 | 280 L | 240L | 225L | 256L | NQ | 252L | 270L | yes
| Temperature | Range: RT +10 to ~ 250 (℃) | C | NQ | yes | yes |  | RT +10 to ~ 250 (℃) | 80-250 | Better, RT+10C to 300C | NQ |  | RT +10 to ~ 300 (℃) | yes
|  | Accuracy: ±0.5 (℃) or better | C | NQ | yes | yes | 0.8 |  | ±0.1 | Yes | NQ |  |  | yes
|  | Temp. Uniformity: ±2(℃) or better | C | NQ | ±2.5 | yes | yes | yes, ±2.5% | No | Yes | NQ | ±2.5 | yes, ±2.5% | yes
| Circulation | Forced air convection | C | NQ | yes | yes | yes | yes | yes | Yes | NQ | yes | yes | yes
| Timer(min) | 1-9999 or continuous | C | NQ | yes | yes | yes | yes | yes | 1 min to 99 days | NQ | yes | yes | yes
| Shelves | Number of shelves: 3 | C | NQ | yes | yes | yes | yes | yes | Yes | NQ | yes | yes | yes
|  | Shelf material: Stainless steel- noncorrosive | C | NQ | yes | yes | yes | yes | yes | Yes | NQ | yes | yes | yes
|  | Adjustable/removable wire shelving | C | NQ | yes | yes | yes | yes | yes | Yes | NQ | yes | yes | yes
| Interior and exterior Material | Interior: Stainless steel- 304 grade Exterior: Powder Coated steel | C | NQ | Stainless steel | yes | yes | No, 201ss, externa electrogalvanized steal | yes | Yes | NQ | Interior: mirror SS, Exterior: Coated steel | yes | yes
| Dimension(W*D*H) | 850 x 880 x 1000 (mm) Do not exceed | C | NQ | 640*680*970 | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | 650*800*1130 | yes | 824 X 1183 X 684mm | NQ | 880*730*950 | 780*705*720 Q: Clarify the correct dimension of the quoted model "BJPX-HGZ270". The dimensions mentioned in the brochure for the quoted model is 900*890*880 (mm). A: Yes, The external size for BJPX-HGZ270 is 900*890*880mm. | 830*875*940
| Weight (kg) | Please mention |  | NQ | 70kg | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | 85kg | yes | Yes | NQ | 110Kg | 60kg | yes
| Power supply | 220/240 V/50 Hz | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | yes | Yes | NQ | yes | yes | yes
| Noice | <60 dB | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | not mentioned | yes |  | Yes | NQ | yes |  | yes
| Safety | Leakage protector and Over- temperature alarm | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | yes | Yes | NQ | yes | yes, over tem[erature alarm | yes
| Digital Display | LCD display | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | yes | Yes | NQ | yes | yes | yes
| Temperature and time controller | PID | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes, PMMA II operating system | yes | Yes | NQ | yes | yes | yes
| Documentation | 1. Manufacture product catalogues with the requested information 2. Manufacture authorization for the bidder to represent them 3. Bidder should have a good after-sales facility with trained technicians and parts available locally to carry out any service/ repair immediately 4. Local Agent should submit a letter from the manufacturer confirming spare parts availability for 5 years. | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | yes | Yes | NQ | yes | yes | 3 years Q: Would you be able to extend the spare parts availability for 5 years for the same quoted price. (Mentioned as 3 years) A: We confirm that we can provide spare parts availability for a period of five (5) years
| Warranty | Local Agent should provide a warranty for 3 years for the instrument and 2 years or more of the warranty period should be from the manufacturer. | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | No 2 years | Yes | NQ | yes | yes | 1 year Q: Warranty period is mentioned as 1 year. would you be able to extend the period to provide a warranty for 3 years for the instrument (2 years or more of the warranty period should be from the manufacturer) for the same quoted price. A: We confirm a two (2) year warranty, at the same quoted price.
| Calibration Service | Should be able to provide annual calibration service free of charge during the warranty period | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | yes | Yes | NQ | yes | yes | yes
| Spare parts availability | Should be available spare parts for the Next 5 Years | C | NQ | yes | Q: Second page of the specification sheet for item 02 not attached in bid document A: Attached in reply | yes | yes | yes | Yes | NQ | yes | yes | 3 years Q: Would you be able to extend the spare parts availability for 5 years for the same quoted price. (Mentioned as 3 years) A: We confirm that we can provide spare parts availability for a period of five (5) years

## 3
[8 empty] | 267000 | [5 empty] | 173508
Item No | Specification | Requirement | Priority | Delmege | research instrument | P&T | Microtech | bhoomitech | Biomedite | Hemsons | premiuM | cmd | Analytical | avon
3 | Equipment | High-precision digital thermometer |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
| Make | (Please specify) | C | NQ | NQ | NQ | NQ | U-therm international (h.k) limited | NQ | NQ | NQ | NQ | NQ | Extech
| Model and Model year | (Please specify) | C | NQ | NQ | NQ | NQ | CX-WDJ200E/2025 | NQ | NQ | NQ | NQ | NQ | EAIIA
| Country of origin | (Please specify) | C | NQ | NQ | NQ | NQ | China | NQ | NQ | NQ | NQ | NQ | USA
| General | A high-precision digital thermometer with external detachable probe designed for laboratory applications, ensuring accurate and reliable temperature measurements in Organic chemistry, biopharmaceutical and research environments. | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | The device should be suitable for monitoring temperature in various laboratory processes. |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
| Power requirement | Battery Type: should provide rechargeable batteries | C | NQ | NQ | NQ | NQ | power supply battery-powered (1AAA battery) | NQ | NQ | NQ | NQ | NQ | yes
|  | Battery Life: Up to 1000 hours of continuous use |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | approx 110h with carbon zinc battery
Voltage: 3V DC or better | bar: NQ to NQ; NQ to NQ
| Temperature Range: | 1.-100°C to +300°C | C | NQ | NQ | NQ | NQ | 50-300 | NQ | NQ | NQ | NQ | NQ | -50 - 1300
|  |  |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
|  | 2.Resolution: 0.01 °C below 200 °C and 0.1 °C above 200 °C |  | NQ | NQ | NQ | NQ | 0.1C | NQ | NQ | NQ | NQ | NQ | yes
|  | 3.Accuracy: ± (0.1% + 0.2°C) below 200°, ±(0.15% + 0.5°C) above 200° | C | NQ | NQ | NQ | NQ | 0.1C | NQ | NQ | NQ | NQ | NQ | ± 0.3% + 2F, ± 0.3% + 1C
|  | 4.Response Time: ≤3 seconds | C | NQ | NQ | NQ | NQ | 5secs | NQ | NQ | NQ | NQ | NQ
|  | 5.Measurement Units: Celsius (°C) and Fahrenheit (°F) | C | NQ | NQ | NQ | NQ | only celcius | NQ | NQ | NQ | NQ | NQ | yes
|  | 6.Sensor Type: Thermocouple or better | C | NQ | NQ | NQ | NQ | Thermocouple | NQ | NQ | NQ | NQ | NQ | yes
| Display & Controls | 1.Display Type: large LCD with backlight or better | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
|  | 2.Control Type: Push buttons for unit selection, hold function, and calibration or better | C | NQ | NQ | NQ | NQ | No control,only display the temper | NQ | NQ | NQ | NQ | NQ | yes
|  | 3.Display should indicate low-battery and probe condition | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | 4.Data Hold Function | C | NQ | NQ | NQ | NQ | NO | NQ | NQ | NQ | NQ | NQ | yes
| Build & Material | Body Material: excellent material resistant to chemical and heat | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
|  | Probe Material: Stainless steel (304 or 316 grade) or better suitable than stainless steel |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | type k probe
|  |  |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
|  | Cable Length: 1m | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | will inform later
|  | IP Rating: IP65 (dust and water-resistant) |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | Not Mentioned
| Accessories | 1.Temperature Probe (detachable) – 1 included | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
Should be provided with rechargeable batteries and a charger | bar: NQ to NQ; NQ to NQ
|  | 2.Protective Carry Case | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | 3. Calibration Certificate | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | 4. User Manual & Warranty Card | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
| Safety Features: | 1.Over-temperature warning | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | 2.Low battery indicator |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | 3.Anti-slip grip for secure handling | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  | 4. Shock-resistant casing | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
|  |  | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ
| Application Areas: | Pharmaceutical and biopharma R&D |  | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
Chemical and analytical laboratories | bar: NQ to NQ; NQ to NQ
Food and beverage industry quality control | bar: NQ to NQ; NQ to NQ
Environmental monitoring | bar: NQ to NQ; NQ to NQ
| Documentation | Manufacture product catalogues with the requested information | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | yes
| Standards | Certifications/Compliance- ISO 17025/ A2LA/ NIST | C | NQ | NQ | NQ | NQ | NO | NQ | NQ | NQ | NQ | NQ | NIST EN61326-1 2013
| Warranty | Local Agent should provide a warranty for the instrument from the manufacturer for 2 years or more. | C | NQ | NQ | NQ | NQ |  | NQ | NQ | NQ | NQ | NQ | 1 year

## 4
[5 empty] | 1227000 |  | 1355000 | 1477000 | [5 empty] | 552844
Item No | Specification | Requirement | Priority | Delmege | research instrument | P&T | Microtech | bhoomitech | Biomedite | Hemsons | premiuM | cmd | Analytical | avon
4 | Equipment | High-Speed Homogenizer |  | NQ |  | NQ |  |  | NQ | NQ | NQ | NQ | NQ
| Make | (Please specify) |  | NQ | DLab | NQ | wiggens | U-therm international (h.k) limited | NQ | NQ | NQ | NQ | NQ | wiggens
| Model and Model year | (Please specify) |  | NQ | D500/2024 | NQ | D 500 pro | D-500/2025 | NQ | NQ | NQ | NQ | NQ | D-500 package 01
| Country of origin | (Please specify) |  | NQ | China | NQ | China | China | NQ | NQ | NQ | NQ | NQ | China
| General Overview | Function: Used for emulsification, dispersion, and homogenization of liquid-liquid or solid-liquid mixtures at high speeds. |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Material | Stainless steel (SS 316) for rotor-stator components | C | NQ | yes | NQ | yes | 316ss | NQ | NQ | NQ | NQ | NQ | yes
| Applications | Suitable for pharmaceutical, biotechnology, food, cosmetics, and chemical laboratories |  | NQ | yes | NQ | yes | yes | NQ | NQ | NQ | NQ | NQ | yes
| Motor Unit | High-performance brushless or AC motor. | C | NQ | AC carbon brush motor | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Power rating: approx. 500 W or better | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Speed range: 5000 to 35,000 RPM (variable speed control) or better suitable to use for emulsification, dispersion, and homogenization of liquid-liquid or solid-liquid mixtures at high speeds. | C | NQ | 10000-30000 | NQ | 500 - 30000 | 10000-30000rpm | NQ | NQ | NQ | NQ | NQ | 10000-30000rpm
|  | Digital display for speed monitoring |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Overload protection for motor safety. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Rotor-Stator Assembly | Rotor diameters (6 mm, 10 mm, 20 mm, 30 mm or better suitable four components to full fill the other requirements in the specification sheet). | C | NQ | 15mm for 10-5000ml, solid to liquid mix, 15 cover volume upto 20L | NQ | Rotor. 15, 4, 9, 23 Stator. 20,5,10,30 | sample processing volume range 10-40000mL | NQ | NQ | NQ | NQ | NQ | yes, 15mm,23mm
|  | High shear forces enable efficient homogenization. |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Easy to dismantle for cleaning and maintenance. |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Homogenizing Shaft | Made of stainless steel or PTFE-coated material or better excellent for chemical resistance. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Immersion depth: 40 mm to 170 mm or better | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Compatible with various rotor-stator sizes |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Designed for easy attachment to the motor unit. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Speed Control Mechanism | Manual or digital speed adjustment. | C | NQ | manual | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Stepless variable speed regulation for precise control. |  | NQ | 6 stage | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Digital models include an LED or LCD display. | C | NQ | manual | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Sample Volume Capacity | Suitable for a wide range of sample volumes: | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | • 500 mL – 10 L or better. |  | NQ | 10-40L | NQ | 10 mL - 40 L |  | NQ | NQ | NQ | NQ | NQ | yes, 10ml-40L
| Cooling System (Optional) | Air-cooled motor housing for prolonged operation. |  | NQ | yes | NQ | NO [not compulsory] | NO | NQ | NQ | NQ | NQ | NQ | yes
|  | Prevents overheating during continuous usage. |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Performance Specifications | Capable of processing viscosities up to 10,000 cP. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Homogenization efficiency: Reduces particle size to below 1 µm. |  | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Noise level: ≤ 72 dB for quiet operation. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Operates in a temperature range of 5°C to 40°C. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  |  | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ
|  |  | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ
| Safety Considerations | Overload protection prevents motor burnout. | C | NQ | yes | NQ | yes | yes | NQ | NQ | NQ | NQ | NQ | yes
|  | Secure mounting stand to ensure stability during operation. | C | NQ | yes | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Safety interlock system | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  | Compliance with CE, and ISO standards. |  | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
| Warranty and Maintenance | Warranty period: minimum 3 years for the whole system, | C | NQ | yes, 3 years | NQ | yes | yes | NQ | NQ | NQ | NQ | NQ | 2 years
|  | Maintenance of the equipment should be every 06 months free of charge during the warranty period |  | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ | yes
|  |  | C | NQ |  | NQ | yes |  | NQ | NQ | NQ | NQ | NQ
| Documentation | Manufacture product catalogues with the requested information | C | NQ | yes | NQ | yes | yes | NQ | NQ | NQ | NQ | NQ | yes

## 5
All specs match. second lowest
[10 empty] | 7406775 |  | 6614992 | 4734800 | 4591665
Item No | Specification | Requirement | Priority | Delmege | research instrument | P&T | Microtech | bhoomitech | Biomedite | Hemsons | premiuM | cmd | Analytical | avon
5 | Equipment | Gel documentation system |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ
| Make | please specify | C | NQ | NQ | NQ | NQ | NQ | NQ | Thistle Scientific | NQ | Infitek | Vilber | uvtec
| Model | please specify | C | NQ | NQ | NQ | NQ | NQ | NQ | gelPRO-302E | NQ | GEP-GD600MF | Quantum cx5 | Platinum Q9
| Country of origin | please specify | C | NQ | NQ | NQ | NQ | NQ | NQ | UK | NQ | China | France | France
| General Usage | It should be compatible with Medical Diagnostic Laboratory use | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ
|  | Should submit relevant proof documents for medical diagnostic laboratory with quotation. |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ | ISO 13485:2016 | Q: Please submit relevant proof documents for medical diagnostic laboratory usage (Mentioned as attached in the specification sheet but the attachment was not available). A: This product has not IVD certification | Q: Please submit relevant proof documents for medical diagnostic laboratory usage. A: Please find the proof document attached for your reference.
|  |  | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ
| Camera & Optics | · Scientific grade camera 05 megapixels or better resolution, which can be extendable later | C | NQ | NQ | NQ | NQ | NQ | NQ | 05 megapixel | NQ | 9MP | yes, 5 mp | 10Mp
|  | · Auto-exposure and auto-focus |  | NQ | NQ | NQ | NQ | NQ | NQ | through genePIX software | NQ |  | yes | yes
|  | · Passive cooling | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes, relies on passive air circulation and CMOS camera used in the system do not generated considerable amount of heat | NQ |  | yes | yes
|  | · 16-bit – more than 60000 grey levels | C | NQ | NQ | NQ | NQ | NQ | NQ | Better, 65,536 grey scale | NQ | 0-65535 | yes | 65535 grey level
|  | · USB-3 interface | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
|  | · Motorized zoom lens with feedback | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
|  | · Field of view: 26 x 21cms or better | C | NQ | NQ | NQ | NQ | NQ | NQ | Better view upto 25 x 30cm | NQ | imaging area 20*20 | yes | yes
|  |  | C | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ
| Hardware | Smart Darkroom technology: | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
|  | • Software control of the lighting |  | NQ | NQ | NQ | NQ | NQ | NQ | Provided quote CS-WLC UV to white converter screen | NQ |  |  | yes
|  | • White light led panels with automatic intensity |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
|  | adjustment |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
|  | • UV cut-off filter and filter wheel with 7 or better positions |  | NQ | NQ | NQ | NQ | NQ | NQ | 7 position filter wheel with UV filter | NQ |  |  | yes
|  | • Slide-out PadBox with automatic recognition of the Application Pad inserted |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ | NO |  | yes
|  | Steel and stainless steel darkroom for long lasting robustness. |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
|  | Wide access door with UV safety shut-off |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
| Laptop | Processor: Minimum Intel Core i5 13th Gen or better RAM: Minimum 8 GB of RAM to support multitasking and efficient software operation. |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes, Quote locally | NQ |  | yes | yes
|  | Display: 14-inch or larger Full HD (1920x1080) display for clear visibility of data and analysis results. |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
|  | Connectivity: At least two USB 3.0 ports, one HDMI output for connecting to external monitors or projectors, and a high-speed Ethernet port or Wi-Fi for data transfer. |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
|  | Battery Life: Minimum 8 hours of battery life for portability during laboratory |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
|  | Ports and Expansion: At least 1 USB 3.0 or USB-C port, and an HDMI output, for easy connection to external devices and data input. |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
|  | Software Compatibility: Original software must be pre-installed or able to support the software required to gel the documentation system (Including original Windows 11 pro software) |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
| Software | Free software for image acquisition with full GLP compliance. | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes, genePIX software provided | NQ |  | yes | yes
|  | Molecular weight calculation, band |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
|  | Quantification, colony counting, distance calculation, text annotation and image enhancement included. |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ |  |  | yes
| Applications | DNA,RNA, protein gels. Fluorescence and colorimetric stain imaging | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
|  | Ethidium bromide, Sybr-Safe, Sybr-Green, Gel-Red, Gel-Green, Sybr-Gold, GFP, Pro-Q Emerald, Sypro ruby, FITC, DAPI |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
|  | Colorimetric stained protein gels, X-Ray film, autorads, SSCP gels, colony dish and flask imaging and other EPI white light applications |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | all listed application available, except xray film imaging
|  | Coomassie blue, Silver stain, Ponceau S Red, Copper stain |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
| Power Supply | Should be operated with mains power supply 240V; 50Hz | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
| Warranty | Minimum 3 year of warranty period for the entire system, Maintenance of the equipment should be every 6 months free of charge during the warranty period | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes, 03-year warranty | NQ |  |  | yes
| Other | Should have a good after sales service. Please attach details of after sales service personnel, their qualification, and address of after sales division etc. | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
|  |  |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ
|  | Documentary evidence of requested specifications should be provided by marking on manufacturer's literature. |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
|  |  |  | NQ | NQ | NQ | NQ | NQ | NQ |  | NQ
|  | A letter of authorization is required from the manufacturer. |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
| Brand name | Brand name should be embedded on the equipment as a permanent mark | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
| Documentation | User manual and other relevant documentations should be provided | C | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  | yes | yes
|  | Should mention catalogue number, manufacturer and authorized dealer in the quotation |  | NQ | NQ | NQ | NQ | NQ | NQ | Yes | NQ |  |  | yes
ISO 9001:2015
Q: Can the device be upgraded within Sri Lanka? A: As per confirmation received from the supplier, camera change can be done locally by a qualified engineer. supplier can be available online if the support is required

## Minutes
Date | Session Time | Attendees Present (List Initials) | Tender Package ID | Equipment/Specification Section | Compliance Status (Select: TRUE/FALSE) | Key Discussion Points/Decisions | Clarification Action Type (Select: Request Issued, Clarification Received, Decision/Closed) | Clarification ID (e.g., Q-YYYY-XXX) | Clarification Status (Select: TRUE/FALSE) | Action Items & Owner | Next Steps/Agenda Item | Document Reference/Link
4 no items match
Hemsons | item 1 | request NMRA/ IVD

## Clarifications
Item | Item Name | Supplier | Clarification/Action Required | Status | Date Sent | Reply
Item 01 | Refrigerated Microcentrifuge | P&T | Please re-attach CE-IVD, NMRA certificate (Mentioned as attached in the specification sheet but the attachment was not available) | Pending | 2025-12-23
Item 01 | Refrigerated Microcentrifuge | Hemsons | Please re-attach CE-IVD, NMRA certificate (Mentioned as attached in the specification sheet but the attachment was not available) | Pending | 2025-12-23
Item 01 | Refrigerated Microcentrifuge | Premium | Please re-attach CE-IVD, NMRA certificate (Mentioned as attached in the specification sheet but the attachment was not available) | Pending | 2025-12-23
Item 02 | Forced Air-Drying Oven | P&T | Second page of the specification sheet not attached (After "Interior and exterior Material") | Pending | 2025-12-23
Item 02 | Forced Air-Drying Oven | Analytical | Clarify the correct dimension of the quoted model "BJPX-HGZ270". The dimensions mentioned in the brochure for the quoted model is 900*890*880 (mm). | Pending | 2025-12-23
Item 02 | Forced Air-Drying Oven | Avon | Would you be able to extend the spare parts availability for 5 years for the same quoted price. (Mentioned as 3 years) | Pending | 2025-12-23
Item 02 | Forced Air-Drying Oven | Avon | Warranty period is mentioned as 1 year. Would you be able to extend the period to provide a warranty for 3 years for the instrument (2 years or more of the warranty period should be from the manufacturer) for the same quoted price. | Pending | 2025-12-23
Item 05 | Gel documentation system | Avon Pharmo Chem Pvt Ltd | Please submit relevant proof documents for medical diagnostic laboratory usage (Mentioned as attached in the specification sheet but the attachment was not available). | Pending | 2025-12-23
Item 05 | Gel documentation system | Avon Pharmo Chem Pvt Ltd | Can the device be upgraded within Sri Lanka? | Pending | 2025-12-23
Item 05 | Gel documentation system | Analytical Instruments Tech Pvt Ltd | Please submit relevant proof documents for medical diagnostic laboratory usage. | Pending | 2025-12-23