# BID EVALUATION REPORT

## FOR

Procurement of Laboratory Equipment for the Sri Lanka Institute of Biotechnology Pvt. Ltd (SLIBTEC), Phase I Building, Mahenwatta, Thalagala Rd, Pitipana, Homagama

File No: SLIBTEC/PROC/211-1/NCB/G&S/Procurement of Laboratory Equipment Supply, Installation and Commissioning

October 2025

## INTRODUCTION

- General

SLIBTEC appointed the committees listed below to execute the procurement activities related to the procurement of laboratory equipment for the Sri Lanka Institute of Biotechnology Pvt. Ltd (SLIBTEC), Phase I Building, Mahenwatta, Thalagala Rd, Pitipana, Homagama.

Department Procurement Committee appointed on 06th August 2025

- Mrs Amali Ranasinghe, Chief Operating Officer, SLIBTEC - Chairperson

- Mrs M M G K Meegahakotuwa, Director General (Planning), Ministry of Science and Technology - Member

- Mr D N Alles, Assistant Director, Department of Public Enterprises - Member

Bid Evaluation Committee appointed on 02nd May 2025

  - Dr Jayantha Senanayaka (Research Professor,SLIBTEC - Chairman

- Snr Professor B.G.D.N.K de Silva (Snr Professor, USJ) – Member

- Dr Dananjaya Perera (Research Scientist, SLIBTEC) – Member

- Dr Peshala Gunasekara (Research Scientist, SLIBTEC) – Member

- Dr Chamith Hewawaduge (Research Scientist, SLIBTEC) – Member

- Dr Duwini Padukkage (Research Scientist, SLIBTEC) – Member

- Dr Chalitha Somarathne (Research Scientist, SLIBTEC) – Member

- Dr Lakmal Gonawala (Research Scientist, SLIBTEC) – Member

- Dr Nalaka Wijekoon (Research Scientist, SLIBTEC) – Member

- Miss P.G.Sachini Raveesha (Quality Executive, SLIBTEC) – Member

- Mr W.G.K Anuradha (Research Associate, SLIBTEC) – Member

- Miss Chathurika Elvitigala (Research Associate, SLIBTEC) – Member

- Mr Nuwan Medawaththa (Research Assistant, SLIBTEC) - Secretary

Quotation/bid opening committee

- Mr. Tharindu Madushankha - Chairperson

- Dr Dananjaya Perera – Member

- Ms Tharangika Bowange – Member

- Mr Navoda Balasooriya – Member

## Package Summary

  - File Number – SLIBTEC/PROC/211-1/NCB/G&S/Procurement of Laboratory Equipment Supply, Installation and Commissioning

  - Package Name- Procurement of Laboratory Equipment Sri Lanka Institute of Biotechnology Pvt. Ltd (SLIBTEC), Phase I Building, Mahenwatta, Thalagala Rd, Pitipana, Homagama.

  - Review Requirement – Prior

  - Procurement Method – National Shopping Procedure

This Procurement includes the following items:

| # Sn No. | # Item | # Quantity |
| --- | --- | --- |
| # 1 | # Refrigerated Microcentrifuge | # 1 Unit |
| # 2 | # Forced Air-Drying Oven | # 1 Unit |
| # 3 | # High-precision digital thermometer | 1 Unit |
| # 4 | # High-Speed Homogenizer | 1 Unit |
| # 5 | # Gel documentation system with a computer | 1 Unit |

## Process Summary

## BEC Appointed on 02nd May 2025

## DPC Meeting held on 08th October 2025 (Budget and Method Approved)

## Bid Documents Checked and Signed by BEC on 10th November 2025

## Bid Documents approved by DPC Members on 14th November 2025

## Bid Documents share to suppliers on 15th November 2025

## Pre-bid Meeting was conducted on 18th November 2025 at 10.15 a.m.

## Bid Opening was conducted on 26th November 2025 at 10.30 a.m.

## Bid Evaluation was conducted from 26th November – 24th December 2025

| Supplier | Date Submitted |
| --- | --- |
| Delmege Forsyth & Co. Ltd. |  |
| Research Instruments Pvt Ltd |  |
| P & T Trading Pvt Ltd |  |
| Microtech Biological Pvt Ltd |  |
| Bhoomi Tech Pvt Ltd |  |
| Biomedite Pvt Ltd |  |
| Hemsons International Pvt Ltd |  |
| Premium International Pvt Ltd |  |
| Commercial Marketing & Distributors Pvt Ltd |  |
| Analytical Instruments Tech Pvt Ltd |  |
| Avon Pharmo Chem Pvt Ltd |  |

- Bid opening

| Bidders | No of items quoted | Bid Opening Price (Rs) (without VAT) |
| --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | 1 | 5,000,000.00 |
| Research Instruments Pvt Ltd | 3 | 5,077,700.00 |
| P & T Trading Pvt Ltd | 2 | 3,247,500.00 |
| Microtech Biological Pvt Ltd | 3 | 5,250,000.00 |
| Bhoomi Tech Pvt Ltd | 4 | 3,711,250.00 |
| Biomedite Pvt Ltd | 2 | 2,763,000.00 |
| Hemsons International Pvt Ltd | 3 | 14,433,670.00 |
| Premium International Pvt Ltd | 1 | 3,350,000.00 |
| Commercial Marketing & Distributors Pvt Ltd | 3 | 9,239,324.00 |
| Analytical Instruments Tech Pvt Ltd | 3 | 7,434,000.00 |
| Avon Pharmo Chem Pvt Ltd | 5 | 12,609,598.00 |

- Evaluation of Quotations

5.1 Completeness of submitted quotations

| Suppliers | Form of quotation is signed | No of quotations submitted | Validity of the offer (requirement in days) | Any other observation |
| --- | --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Yes |  |  | No |
| Research Instruments Pvt Ltd | Yes |  |  | No |
| P & T Trading Pvt Ltd | Yes |  |  | No |
| Microtech Biological Pvt Ltd | Yes |  |  | No |
| Bhoomi Tech Pvt Ltd | Yes |  |  | No |
| Biomedite Pvt Ltd | Yes |  |  | No |
| Hemsons International Pvt Ltd | Yes |  |  | No |
| Premium International Pvt Ltd | Yes |  |  | No |
| Commercial Marketing & Distributors Pvt Ltd | Yes |  |  | No |
| Analytical Instruments Tech Pvt Ltd | Yes |  |  | No |
| Avon Pharmo Chem Pvt Ltd | Yes |  |  | No |

5.2 Eligibility of Bidders

| Suppliers | Registered company in Sri Lanka | Potential conflict of interest | Any other observations |
| --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Yes | No | No |
| Research Instruments Pvt Ltd | Yes | No | No |
| P & T Trading Pvt Ltd | Yes | No | No |
| Microtech Biological Pvt Ltd | Yes | No | No |
| Bhoomi Tech Pvt Ltd | Yes | No | No |
| Biomedite Pvt Ltd | Yes | No | No |
| Hemsons International Pvt Ltd | Yes | No | No |
| Premium International Pvt Ltd | Yes | No | No |
| Commercial Marketing & Distributors Pvt Ltd | Yes | No | No |
| Analytical Instruments Tech Pvt Ltd | Yes | No | No |
| Avon Pharmo Chem Pvt Ltd | Yes | No | No |

5.3 Contractual Experience

| Supplier | Work Experience (Covered one contract over the last 3 years) |
| --- | --- |
| Delmege Forsyth & Co. Ltd. | Yes |
| Research Instruments Pvt Ltd | Yes |
| P & T Trading Pvt Ltd | Yes |
| Microtech Biological Pvt Ltd | Yes |
| Bhoomi Tech Pvt Ltd | Yes |
| Biomedite Pvt Ltd | Yes |
| Hemsons International Pvt Ltd | Yes |
| Premium International Pvt Ltd | Yes |
| Commercial Marketing & Distributors Pvt Ltd | Yes |
| Analytical Instruments Tech Pvt Ltd | Yes |
| Avon Pharmo Chem Pvt Ltd | Yes |

5.4 Compliance to Technical Specifications

From the eligible bids, only the bids that satisfied the required specifications were considered for price evaluation for the Procurement of Laboratory Equipment for the Sri Lanka Institute of Biotechnology Pvt. Ltd (SLIBTEC), Phase I Building, Mahenwatta, Thalagala Rd, Pitipana, Homagama.

  - Refrigerated Microcentrifuge

| Bidder | Compliance to Specifications | Make and Model | Manufacture Country |
| --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Complies with all the specifications | Thermo Scientific - Multifuge X1R Pro/ 2021 | Germany |
| Research Instruments Pvt Ltd | Not complying with the specifications - No NMRA registration or CE-IVD certification submitted for diagnostic use. | Dlab DG1616R | China |
| P & T Trading Pvt Ltd | Not complying with the specifications - Submitted certifications are neither NMRA registration nor CE-IVD certification. | YINGTAI TGL18 | China |
| Microtech Biological Pvt Ltd | Not complying with the specifications - Only IVD certification submitted. | ortoalresa Digicen 22R/2025 | Spain |
| Bhoomi Tech Pvt Ltd | Not complying with the specifications - Only CE certification submitted. - Warranty requirement not met. UPS system warranty provided for 1 year only; minimum required warranty is 2 years for both the instrument and UPS. | SciTEK CFG-TI5HR/2025 | China |
| Biomedite Pvt Ltd | Not complying with the specifications - No NMRA registration or CE-IVD certification submitted. - Acceleration/deceleration profiles limited to 7; minimum requirement is 9 profiles with pulse-free operation. | shuke TGL-1650,2025 | China |
| Hemsons International Pvt Ltd | Not complying with the specifications - CE certification submitted does not include IVD scope. | Ohaus FC5718RIVD | USA |
| Premium International Pvt Ltd | Not complying with the specifications - Submitted certifications are neither NMRA registration nor CE-IVD certification. | Hettich Universal 320R | Germany |
| Commercial Marketing & Distributors Pvt Ltd | Not complying with the specifications - Only CE certification submitted. | Infitek CFGR-B21B | China |
| Analytical Instruments Tech Pvt Ltd | Not complying with the specifications - Maximum Speed with the given rotor/ rpm ≥ 15,000 rpm : 14800rpm | ThermoScientific Sorvall legend Micro 21R | China |
| Avon Pharmo Chem Pvt Ltd | Not complying with the specifications - Maximum Speed with the given rotor/ rpm ≥ 15,000 rpm : 14000rpm - Only CE-IVD certification submitted | Eppendorf 5804R/5805000327 | Germany |

Note: As per the Pre-Bid Meeting Minutes, priority will be given to equipment supported by both NMRA registration and CE-IVD certification (Annexure 01)

  - Forced Air-Drying Oven

| Bidder | Compliance to Specifications | Make and Model | Manufacture Country |
| --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Not Quoted |  |  |
| Research Instruments Pvt Ltd | Not complying with the specifications - The required minimum capacity of 250 L is not met; the offered capacity is 140 L. | Infitek DOF-v140A | China |
| P & T Trading Pvt Ltd | Not complying with the specifications - The required minimum capacity of 250 L is not met; the offered capacity is 240 L. | Bio Base BOV-V225FI | China |
| Microtech Biological Pvt Ltd | Not complying with the specifications - The specified temperature accuracy of ±0.5 °C or better is not achieved; the declared accuracy is ±0.8 °C. | Boxun BGZ-306 | China |
| Bhoomi Tech Pvt Ltd | Not complying with the specifications - The required minimum capacity of 250 L is not met; the offered capacity is 240 L. - The required construction materials are not met; the interior is not stainless steel grade 304 and the exterior is electro-galvanized steel instead of powder-coated steel. | SCITEK FDO-240A | China |
| Biomedite Pvt Ltd | Not complying with the specifications - The required minimum capacity of 250 L is not met; the offered capacity is 225 L. - The specified temperature uniformity of ±2 °C or better is not met; no compliant uniformity performance is provided. - The warranty requirement is not fulfilled; a minimum of three (3) years warranty with at least two (2) years from the manufacturer is required, whereas only a two (2) year warranty is offered. | China-Zenith LAB PPH-240A | China |
| Hemsons International Pvt Ltd | Complies with all the specifications | Memmert UF260 | Germany |
| Premium International Pvt Ltd | Not Quoted |  |  |
| Commercial Marketing & Distributors Pvt Ltd | Not complying with the specifications - The required interior material of stainless steel grade 304 and exterior material of powder-coated steel are not met; the offered unit has a mirror stainless steel interior and a coated steel exterior. | Infitek DOF-H250B | China |
| Analytical Instruments Tech Pvt Ltd | Not complying with the specifications - The maximum allowable dimensions of 850 × 880 × 1000 mm are exceeded; the offered dimensions are 900 × 890 × 880 mm. | biobase BJPX-HGZ270 | China |
| Avon Pharmo Chem Pvt Ltd | Complies with all the specifications | LK lab LO-FS250 | Korea |

  - High-precision digital thermometer

| Bidder | Compliance to Specifications | Make and Model | Manufacture Country |
| --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Not Quoted |  |  |
| Research Instruments Pvt Ltd | Not Quoted |  |  |
| P & T Trading Pvt Ltd | Not Quoted |  |  |
| Microtech Biological Pvt Ltd | Not Quoted |  |  |
| Bhoomi Tech Pvt Ltd | Not complying with the specifications - The specified temperature range of –100 °C to +300 °C is not met; the offered range is 50 °C to 300 °C. - The required temperature resolution of 0.01 °C below 200 °C is not provided; a resolution of 0.1 °C is offered across all ranges. - The response time requirement of ≤ 3 seconds is not met; the declared response time is 5 seconds. - Measurement in Fahrenheit (°F) is not available; the device supports Celsius (°C) only. - Required control functions via push buttons for unit selection, data hold are not provided; the unit has temperature display only. - Compliance with ISO 17025 / A2LA / NIST standards has not been provided. | U-therm international (h.k) limited CX-WDJ200E/2025 | China |
| Biomedite Pvt Ltd | Not Quoted |  |  |
| Hemsons International Pvt Ltd | Not Quoted |  |  |
| Premium International Pvt Ltd | Not Quoted |  |  |
| Commercial Marketing & Distributors Pvt Ltd | Not Quoted |  |  |
| Analytical Instruments Tech Pvt Ltd | Not Quoted |  |  |
| Avon Pharmo Chem Pvt Ltd | Not complying with the specifications - The specified battery life of up to 1000 hours of continuous use is not met; the offered battery life is approximately 110 hours using a carbon zinc battery. - The required temperature accuracy of ±(0.1% + 0.2 °C) below 200 °C and ±(0.15% + 0.5 °C) above 200 °C is not met; the declared accuracy is ±0.3% + 2 °F and ±0.3% + 1 °C. - The required manufacturer’s warranty of a minimum of two (2) years is not met; only a one (1) year warranty is provided. | Extech EAIIA | USA |

  - High-Speed Homogenizer

| Bidder | Compliance to Specifications | Make and Model | Manufacture Country |
| --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Not Quoted |  |  |
| Research Instruments Pvt Ltd | Not complying with the specifications - The required motor unit speed range of 5,000 to 35,000 RPM with variable speed control is not met; the offered speed range is 10,000 to 30,000 RPM. - The required rotor–stator assembly diameters of 6 mm, 10 mm, 20 mm, and 30 mm (or better) are not provided; only a 15 mm rotor is offered, which does not meet the specified range or configuration requirements. | DLab D500/2024 | China |
| P & T Trading Pvt Ltd | Not Quoted |  |  |
| Microtech Biological Pvt Ltd | Complies with all the specifications* | wiggens D 500 pro | China |
| Bhoomi Tech Pvt Ltd | Not complying with the specifications - The required motor unit speed range of 5,000 to 35,000 RPM with variable speed control is not met; the offered speed range is 10,000 to 30,000 RPM. | U-therm international (h.k) limited D-500/2025 | China |
| Biomedite Pvt Ltd | Not Quoted |  |  |
| Hemsons International Pvt Ltd | Not Quoted |  |  |
| Premium International Pvt Ltd | Not Quoted |  |  |
| Commercial Marketing & Distributors Pvt Ltd | Not Quoted |  |  |
| Analytical Instruments Tech Pvt Ltd | Not Quoted |  |  |
| Avon Pharmo Chem Pvt Ltd | Not complying with the specifications - The required motor unit speed range of 5,000 to 35,000 RPM with variable speed control is not met; the offered speed range is 10,000 to 30,000 RPM. - The required rotor–stator assembly diameters of 6 mm, 10 mm, 20 mm, and 30 mm (or better) are not provided; only 15 mm and 23 mm rotor is offered, which does not meet the specified range or configuration requirements. - The required minimum warranty period of three (3) years for the complete system is not met; only a two (2) year warranty is provided. | wiggens D-500 package 01 | China |

*Note - The optional requirement for an air-cooled motor housing for prolonged operation is not provided. This deviation was clarified with the equipment requestor and has been confirmed by the Bid Evaluation Committee (BEC) as acceptable.

  - Gel documentation system with a computer

| Bidder | Compliance to Specifications | Make and Model | Manufacture Country |
| --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | Not Quoted |  |  |
| Research Instruments Pvt Ltd | Not Quoted |  |  |
| P & T Trading Pvt Ltd | Not Quoted |  |  |
| Microtech Biological Pvt Ltd | Not Quoted |  |  |
| Bhoomi Tech Pvt Ltd | Not Quoted |  |  |
| Biomedite Pvt Ltd | Not Quoted |  |  |
| Hemsons International Pvt Ltd | Complies with all the specifications | Thistle Scientific gelPRO-302E | UK |
| Premium International Pvt Ltd | Not Quoted |  |  |
| Commercial Marketing & Distributors Pvt Ltd | Not complying with the specifications - The required slide-out PadBox with automatic recognition of the inserted application pad is not provided. | Infitek GEP-GD600MF | China |
| Analytical Instruments Tech Pvt Ltd | Not complying with the specifications - Relevant proof documents for use in a medical diagnostic laboratory, including IVD certification, were not submitted. | Vilber Quantum cx5 | France |
| Avon Pharmo Chem Pvt Ltd | Complies with all the specifications* | uvtec Platinum Q9 | France |

*Note - The bidder initially indicated spare parts availability for three (3) years and a one (1) year warranty period. Upon clarification, the bidder confirmed that spare parts will be made available for a period of five (5) years at the same quoted price. The bidder further confirmed a two (2) year warranty at the same quoted price; however, this does not fully meet the tender requirement of a three (3) year warranty with a minimum of two (2) years from the manufacturer. The requestor has reviewed the clarifications and confirmed acceptance of this deviation.

- Price Schedule

  - Refrigerated Microcentrifuge

| Bidder | Quoted price per Unit (without VAT) LKR | Arithmetic error Correction | Evaluated price LKR (Total price excluding VAT) | Rank |
| --- | --- | --- | --- | --- |
| Delmege Forsyth & Co. Ltd. | 5,000,000.00 | - | 5,000,000.00 | 1 |

  - Forced Air-Drying Oven

| Bidder | Quoted price per Unit (without VAT) LKR | Arithmetic error Correction | Evaluated price LKR (Total price excluding VAT) | Rank |
| --- | --- | --- | --- | --- |
| Avon Pharmo Chem Pvt Ltd | 1,383,660.00 | - | 1,383,660.00 | 1 |
| Hemsons International Pvt Ltd | 2,332,865.00 | - | 2,332,865.00 | 2 |

  - High-precision digital thermometer

None of the quoted item complies with all the specifications.

  - High-Speed Homogenizer

| Bidder | Quoted price per Unit (without VAT) LKR | Arithmetic error Correction | Evaluated price LKR (Total price excluding VAT) | Rank |
| --- | --- | --- | --- | --- |
| Microtech Biological Pvt Ltd | 1,355,000.00 | - | 1,355,000.00 | 1 |

  - Gel documentation system with a computer

| Bidder | Quoted price per Unit (without VAT) LKR | Arithmetic error Correction | Evaluated price LKR (Total price excluding VAT) | Rank |
| --- | --- | --- | --- | --- |
| Avon Pharmo Chem Pvt Ltd | 4,591,665.00 | - | 4,591,665.00 | 1 |
| Hemsons International Pvt Ltd | 7,406,775.00 | - | 7,406,775.00 | 2 |

7. Recommendation of the TEC members

7.1 Technical Evaluation Committee has recommended the DPC to award the contract to the following suppliers mentioned in the table after the evaluation.

- Refrigerated Microcentrifuge

Bid Evaluation Committee has recommended the DPC to award the contract to Delmege Forsyth & Co. Ltd. being the lowest bidder, whose items were compatible with the required specifications at a total cost of Five Million in Sri Lankan Rupees (LKR 5,000,000.00) without VAT.

| S No | Item | Quantity | Unit Price | Total Price |
| --- | --- | --- | --- | --- |
| 01 | Refrigerated Microcentrifuge | 01 | 5,000,000.00 | 5,000,000.00 |
| Evaluated Total (Without VAT) |  |  |  | 5,000,000.00 |

- Forced Air-Drying Oven

Bid Evaluation Committee has recommended the DPC to award the contract to Avon Pharmo Chem Pvt Ltd being the lowest bidder, whose items were compatible with the required specifications at a total cost of One Million Three Hundred Eighty-Three Thousand Six Hundred Sixty in Sri Lankan Rupees (LKR 1,383,660.00) without VAT.

| S No | Item | Quantity | Unit Price | Total Price |
| --- | --- | --- | --- | --- |
| 02 | Forced Air-Drying Oven | 01 | 1,383,660.00 | 1,383,660.00 |
| Evaluated Total (Without VAT) |  |  |  | 1,383,660.00 |

- High-precision digital thermometer

Bid Evaluation Committee has decided none of the quoted items complies with all the specifications

- High-Speed Homogenizer

Bid Evaluation Committee has recommended the DPC to award the contract to Microtech Biological Pvt Ltd being the lowest bidder, whose items were compatible with the required specifications at a total cost of One Million Three Hundred Fifty-Five Thousand in Sri Lankan Rupees (LKR 1,355,000.00) without VAT.

| S No | Item | Quantity | Unit Price | Total Price |
| --- | --- | --- | --- | --- |
| 04 | High-Speed Homogenizer | 01 | 1,355,000.00 | 1,355,000.00 |
| Evaluated Total (Without VAT) |  |  |  | 1,355,000.00 |

- Gel documentation system with a computer

Bid Evaluation Committee has recommended the DPC to award the contract to Delmege Forsyth & Co. Ltd. being the lowest bidder, whose items were compatible with the required specifications at a total cost of Five Million in Sri Lankan Rupees (LKR 4,591,665.00) without VAT.

| S No | Item | Quantity | Unit Price | Total Price |
| --- | --- | --- | --- | --- |
| 05 | Gel documentation system with a computer | 01 | 4,591,665.00 | 4,591,665.00 |
| Evaluated Total (Without VAT) |  |  |  | 4,591,665.00 |

7.2 Clarifications

●	The committee has requested clarifications related to this procurement process, and enclosed herewith.

●	A copy of requested clarifications from each bidder and their response has been attached with the document (Annexure 02).

7.3 Conclusion

BEC recommends to proceed with procurement as follows.

The Bid Evaluation Committee has recommended the DPC to award One Refrigerated Microcentrifuge to Delmege Forsyth & Co. Ltd. at a total cost of Five Million in Sri Lankan Rupees (LKR 5,000,000.00) without VAT. Total price including VAT Five Million Nine Hundred Thousand in Sri Lankan Rupees (LKR 5,900,000.00).

The Bid Evaluation Committee has recommended the DPC to award One Forced Air-Drying Oven and One Gel documentation system with a computer to Avon Pharmo Chem Pvt Ltd at a total cost of Five Million Nine Hundred Seventy-Five Thousand Three Hundred Twenty-Five in Sri Lankan Rupees. (LKR 5,975,325.00) without VAT. Total price including VAT Seven Million Fifty Thousand Eight Hundred Eighty-Three and Fifty Cents in Sri Lankan Rupees (LKR 7,050,883.50).

The Bid Evaluation Committee has recommended the DPC to award One High-Speed Homogenizer to Microtech Biological Pvt Ltd at a total cost of One Million Three Hundred Fifty-Five Thousand in Sri Lankan Rupees. (LKR 1,355,000.00) without VAT. Total price including VAT One Million Five Hundred Ninety-Eight Thousand Nine Hundred in Sri Lankan Rupees (LKR 1,598,900.00).

| Member | Opinion | Signature | Date |
| --- | --- | --- | --- |
| Dr Jayantha Senanayake |  |  |  |
| Snr.Professor.B.G.D.N.K.de Silva |  |  |  |
| Dr Dananjaya Perera |  |  |  |
| Dr Peshala Gunasekara |  |  |  |
| Dr Duwini Padukkage |  |  |  |
| Dr Chamith Hewawaduge |  |  |  |
| Dr Chalitha Somarathne |  |  |  |
| Dr Lakmal Gonawala |  |  |  |
| Dr Nalaka Wijekoon |  |  |  |
| Miss P.G.Sachini Raveesha |  |  |  |
| Mr W. G. K. Anuradha |  |  |  |
| Miss Chathurika Elvitigala |  |  |  |
| Mr Nuwan Medawaththa |  |  |  |